## Sheet1
ARA Flow Calibration Worksheet
Site Information
Location: | Eugene, OR | Sampler: | N-FRM | Serial No: | 10000
Tech: | John Doe | Flow Std: | FTS-A | Serial No: | 10000
Date: | 2017-01-01 | Temp Std: | FTS-A | Serial No: | 10000
Time: | 0:00 | Pressure Std: | FTS-A | Serial No: | 10000
Site Conditions
Reference Pressure (mmHg) |  | 743 | Reference Temperature (°C) |  | 21.2
Sampler Pressure (mmHg) |  | 744 | Sampler Temperature (°C) |  | 21.1
Barometric Pressure Offset |  | -1 | Temperature Offset: |  | 0.1
Pre-Calibration
Set Flow Rate | (Sampler) Indicated Flow LPM @LTP* | (Reference) Actual Flow LPM @LTP | Percent Difference | Pre-Calibration Constants
Leak Check | 0 |  |  | Slope | 1
16.7 | 16.5 | 16.7 | -1.1976 | Intercept | 0
Prior to Calibration : manually set Slope =1.000 & Intercept = 0.000 in Sampler @
SETUP>SYSTEM SETUP>FLOW RATE
Calibration Information
Set Flow Rate | (Sampler) Indicated Flow LPM @LTP | (Reference) Actual Flow LPM @LTP | Percent Difference | Adjusted Flow | Percent Error
14.5 | 14.49 | 14.62 | -0.88919 | 14.6341 | -0.096222
15.5 | 15.5 | 15.64 | -0.89514 | 15.643 | -0.019359
16.5 | 16.49 | 16.66 | -1.0204 | 16.632 | 0.16802
17.5 | 17.5 | 17.65 | -0.84986 | 17.641 | 0.051172
18.5 | 18.51 | 18.63 | -0.64412 | 18.6499 | -0.10697
Post Cal 16.7 | 16.65 | 16.7 | -0.2994
Slope | 0.99897 |  | * LTP= Local Temperature & Pressure
Intercept | 0.15899
Enter These Calibration Constants Into Sampler @
SETUP>SYSTEM SETUP>FLOW RATE